Journal of the Laser and Health Academy

ISSN 1855-9913 Vol. 2011, No. 1; www.laserandhealth.com

**New Skin Treatment Possibilities with PIANO Mode on an Nd:YAG Laser**

**Matjaz Lukac,1 Zdenko Vizintin, 2 Samo Pirnat,3 Karolj Nemes2**

*1Jozef Stefan Institute, Light and Matter Department, Jamova cesta 39, Ljubljana, Slovenia*

*2Fotona d. d., Stegne 7, Ljubljana, Slovenia*

**ABSTRACT**

Non-ablative skin textural improvement is based on the theory of delivering a suitable packet of photo thermal damage (coagulative and biostimulating) to dermal layers beneath a cooled epidermis. In this paper, a new Nd:YAG laser pulse duration modality is introduced which is perfectly suited for this purpose. This new, super long, PIANO modality, which extends the Nd:YAG pulse durations to the seconds regime, is much longer than the thermal relaxation time of the epidermis or any other skin structures, and does not cause high initial temperature peaks in the epidermis. It is therefore indicated for treatments where overall homogeneous, bulk heating of the dermis is desired. Based on the in-vivo study of the thermal effects of the PIANO mode on human skin, clinical protocol guidelines for laser skin remodeling, skin tightening and laser-assisted wound healing are developed. The PIANO modality may be of benefit also when considering treatments that combine the non-ablative Nd:YAG laser skin remodeling with the Er:YAG laser fractional skin resurfacing.

**Key words:** laser skin remodeling, scar prevention, wound healing, Nd:YAG laser, PIANO mode.

*Article: J. LAHA, Vol. 2011, No.1; pp. 22-32. Received: March 03, 2011; Accepted: March 30, 2011.*

*© Laser and Health Academy. All rights reserved.* Printed in Europe. *www.laserandhealth.com*

**I. INTRODUCTION**

**a) Selective photo-thermolysis**

Since the invention of the laser, many types of laser sources have been studied and used as a means to improve skin appearance [1]. If the clinical objective is to cause selective modifications of a specific tissue structure, the laser wavelength should match the highest absorption of the targeted structure relative to the surrounding tissue. This approach is called “selective photothermolysis” [2]. Typically, however, the wavelengths that are highly absorbed in skin

imperfections are also highly absorbed by non-target structures, for example, melanosomes [3] or hemoglobin-containing RBC [4]. Consequently, these wavelengths do not reach deeper lying targets (such as blood vessels or hair follicles), and can result in excessive damage to the epidermis and other healthy skin structures. The choice of wavelength is thus dictated not only by the need for good absorption of the laser in the treated skin structure, but also by the need to avoid unnecessary damage to the epidermis. For this reason, it is often better to select a laser wavelength that penetrates more deeply into the tissue, and then achieve selective tissue modification by adjusting the laser pulse duration to the thermal relaxation time of the targeted imperfection.

Among all non-ablative laser sources, the Nd:YAG laser holds the most prominent position. This is due to its long wavelength of 1064 nm, which in terms of absorption, lies in an optical window that allows light of this wavelength to penetrate deep into the skin, while its absorption in a target such as a blood vessel or a hair follicle is strong enough to affect the target (See Fig. 1) [5, 6].



Fig. 1: Absorption coefficient in human tissues, μa as a function of laser wavelength. The absorption coefficient is at a minimum with the Nd:YAG laser wavelength (1.064 μm), and at a maximum with the Er:YAG laser wavelength (2.94 μm).

Besides the proper selection of the wavelength in order to illuminate a given tissue structure, the pulse duration must also be considered. Namely, the pulse

22

New Skin Treatment Possibilities with PIANO Mode on an Nd:YAG Laser

duration determines the heat source generated in the tissue since the heat diffusion process leads to temperature gradients in it. During a lengthy laser exposure most of the deposited heat will diffuse away from the target structure, resulting in nonspecific thermal damage to adjacent structures. Conversely, a laser pulse that is significantly shorter than the target thermal relaxation time TRT minimizes the time

mode) [6-7] (See Fig. 4).

300

250

200

**)**

150

available for heat diffusion and confines the heating

**m**

μ

**(**

effect to the target structure, resulting in maximum

**d**

100

temperature difference between the target and 50

adjacent structures (see Fig. 2) [7]. Here, the relaxation time TRT represents the time interval in which the

0

amplitude of a hypothetical temperature rise decreases by approximately a factor of two due to the diffusion of heat into surrounding tissue.

0 1 2 3 4 5 6 7 8 9 10 **Pulse duration tp (ms)**

LASER

LASER

Laser beam

Laser beam

Laser beam

Laser beam

BEAM

BEAM

SKIN

SKIN

target

target

with

with

with

with

Fig. 3: Minimal size (d) of a skin structure that can be selectively heated by a laser pulse of duration tp.

**POWER**

**POWER**

Q-SWITCH

Q-SWITCH

**Megawatts**

**Megawatts**

t ΔT p < TRTtarget

t ΔT p < TRTtarget ΔT

ΔT

ΔT

FRAC3 FRAC3

ΔT

ΔT

ΔT

ΔT

ΔT

ΔT

ΔT

FRAC3

FRAC3

FRAC3

FRAC3

pulse

pulse

pulse

pulse

with

with

with

with

tp > TRTtarget tp > TRTtarget standard

standard

standard

standard

pulse

pulse

pulse

**Kilowatts Kilowatts**

**Watts**

**Watts**

LP

LP

**Nanoseconds Microseconds Miliseconds Seconds Nanoseconds Microseconds Miliseconds Seconds**

**TIME TIME**

pulse

Fig. 2: With short laser pulsewidth (tp < TRTtarget), a small skin imperfection or a hair follicle experiences a high localized temperature increase ΔT, while with a longer laser pulsewidth (tp > TRTtarget), the local temperature increase is much smaller.

Figure 3 shows the dependence of the minimal size d of the skin target that can be selectively heated by a laser pulse of duration tp. Confinement of laser energy within smaller structures requires progressively shorter pulse durations [8].

Using the relatively deeply penetrating Nd:YAG laser wavelength, and targeting skin imperfections by adjusting the laser pulse duration tp to the cooling times of these imperfections, is the paradigm behind most of the Nd:YAG laser aesthetic procedures [6-8]. For this reason, aesthetic Nd:YAG laser systems cover a wide range of laser pulse durations, from nanoseconds (Q-switched mode) [9] to several hundred microseconds (FRAC3 mode) [8], and finally to several tens of milliseconds (LP or Long Pulse

Fig. 4: Wide range of Nd:YAG pulse durations and instantaneous pulse powers.

All of the above pulse duration modes are used for selective photo-thermolysis procedures, with the optimal choice of the pulse duration depending on the size of the treated skin structure. The Q-switch pulses are preferably used for treating tattoos and pigmented lesions [9]. The FRAC3 pulses produce a three

dimensional fractional pattern in the epidermis and dermis, with damage islands located predominantly at the sites of miniscule aging skin imperfections [8].

And the duration of LP pulses is matched with the relaxation times of hair follicles and blood vessels [6- 7].

**b) Homogeneous photo-thermolysis with the new PIANO mode**

Laser treatment is safest when a laser pulsewidth can be chosen that is longer than the thermal relaxation time, TRTepi of the epidermis, but shorter than the TRTtarget of the treated skin target. This

23

New Skin Treatment Possibilities with PIANO Mode on an Nd:YAG Laser

allows selective heat treatment of the target without overheating the epidermis.

A schematic presentation of the influence of laser pulsewidth on the temperature distribution within the skin is shown in Fig. 5 [7].

ΔTepi

skin.

In non-tanned, light skin photo-types, melanosomes are concentrated primarily in a depi =10 μm thick basal layer (typically located 50-100 μm below the skin surface) while in tanned, darker skin phototypes more melanin is distributed throughout the epidermis. As a result, the epidermal TRT can vary from less than 1 ms to over 100 ms [19]. In most

) )

.

.

u

u

.

.

a

a

((

TT

Δ

Δ

|  |
| --- |
|  |
|  |

ΔTepi

a)

a)

cases, however, the thermal relaxation time of the epidermis is on the order of 25-100 ms. Figure 6 shows results of a recent study [7], where epidermal temperature decay curves were measured following an irradiation with a short duration (1 ms) Nd:YAG laser pulse.

0

0

) )

.

.

u

u

.

.

a

a

((

T

T

Δ

Δ

epidermis

epidermis

dermis

dermis

1234 1234 b)

b)

ΔTepi

ΔTepi

**e** 

**e**

**r**

**r**

**u**

**u**

**t**

**t**

**a**

**a**

**r**

**r**

**e**

**e**

**p**

**p**

**i**

**i**

**m**

**m**

**A**

**A**

**p**

**p**

**e**

**e**

**e**

**e**

**t**

**t**

**B**

**B**

**T**

**T**

**d**

**d**

**e**

**e**

Δ

Δ

**C**

**C**

**z**

**z**

**i**

**i**

**l**

**l**

**a**

**a**

**m**

**m**

**r**

**r**

**o**

**o**

**N**

**N**

**Time (ms)**

**Time (ms)**

0

0

1234

1234

Skin depth (mm)

Skin depth (mm)

.

Fig. 6: Skin surface normalized temperature (Tepi) decay curves following a 1ms pulse duration Nd:YAG laser pulse for geographical skin types A (Europe), B (China, Egypt) and C (India, Japan) [7].

Fig. 5: Calculated temperature distribution immediately after a short Nd:YAG laser pulse with tp << TRTepi (a), and after a long duration Nd:YAG laser pulse with tp >> TRTepi (b). For skin types I-II, the temperature increase of epidermis is approximately 3 x higher when short pulse durations are used [7].

For short laser pulsewidths (tp << TRTepi), the distribution of the temperature increase (ΔT) closely follows the absorption characteristics of the epidermis and dermis. Due to the high concentration of melanosomes in the epidermis, laser absorption, and consequently ΔT is highest in the approximately 100 μm thick epidermal layer, while the temperature increase in the dermis is more moderate (Fig. 5a).

For long laser pulsewidths (tp >> TRTepi), the temperature distribution is dominated by the thermal diffusion process. Since there is a sufficient time for the epidermis to share the absorbed heat with dermis through heat diffusion, the temperature distribution does not have a pronounced peak in the epidermal layer (Fig. 5 b). Long pulse durations therefore result in a relatively homogeneous “bulk” heating of the

In Fig. 6, the initial fast decay lasts approximately 100 ms, and is caused by the diffusion mediated heat transfer from epidermis to dermis. This is followed by a slower decay that lasts several seconds. During this slower cooling process, the bulk skin looses heat due to its thermal contact with the ambient air and deeper body tissues.

Therefore, when no specific skin imperfections are targeted, and instead a non-selective heating, or “homogeneous photo-thermolysis” of the bulk dermis is desired, Nd:YAG laser pulsewidths considerably longer than 100 ms are preferable. Longer pulsewidths ensure that the dermis is homogeneously heated up while the epidermis is spared from unnecessary thermal damage.

Most of the presently available aesthetic Nd:YAG laser systems do not operate at pulsewidths much longer than 100 ms. They have been developed for the selective photo-thermolysis of skin imperfections that requires pulsewidths that match relatively short thermal relaxation times of the targeted skin imperfections. There are, however, several important

24

New Skin Treatment Possibilities with PIANO Mode on an Nd:YAG Laser

indications for homogeneous photo-thermolysis that require “super long” laser pulsewidths. These indications include photo-thermal skin remodeling, skin tightening and surgical scar prevention.

Recently, a new Nd:YAG laser pulse duration modality has been developed that is intended specifically for homogeneous, photo-thermolysis treatments [10]. This new, PIANO modality extends the Nd:YAG pulse durations to the seconds regime (0.3 - 60 s), and reduces treatment pulse powers to a less invasive watts regime (up to 80 W) (See Fig. 7).

**POWER**

**POWER**

Q-SWITCH

Q-SWITCH

**Megawatts**

**Megawatts**

E

E

D

D

O

O

 M

 M

FRAC3

the experiments was a Fotona XD-2 810 nm diode laser system.



FRAC3

**Kilowatts**

**Kilowatts**

**Watts**

**Watts**

LP LP

W

W

E

E

N

N

PIANO

PIANO

**TIME**

**TIME**

Fig. 8: Fotona XP Dynamis Nd:YAG laser system with PIANO mode capabilities, used in the study [15].

**Nanoseconds Microseconds Miliseconds Seconds**

**Nanoseconds Microseconds Miliseconds Seconds**

Fig. 7: The new “super long pulse” PIANO mode enables pulse durations in the seconds regime, and exhibits less intensive pulse powers of several Watts.

The duration of the PIANO mode is much longer than the thermal relaxation time of the epidermis. This is therefore the safest mode for reaching deeper lying skin tissues, with minimal thermal effect on the epidermis. The PIANO mode pulsed durations are also longer than the relaxation times of any other skin structures, such as hair follicles or blood vessels. The PIANO modality is thus indicated for overall homogeneous, bulk heating of the dermis, justifying its name (piano: adjective meaning soft or smooth [11]).

In this paper, we report on an in-vivo study of the latest PIANO mode in terms of its thermal effect on human skin. Based on the obtained results and previously published clinical data, clinical protocol guidelines were developed for the following homogeneous photo-thermolysis treatments: a) skin remodeling, b) dual wavelength skin resurfacing, and c) wound healing and scar prevention.

**II. MATERIALS AND METHODS**

The Nd:YAG laser used in the study was a Fotona XP Dynamis laser system with R33 and R34 non contact handpieces with spotsize diameters from 2-20 mm (See Fig. 8). The laser was upgraded to the new PIANO (tp=0.3 - 60 s) mode. The diode laser used in

A thermal imager, model Sagem Matis, operating in the 3-5 μm spectral range, was fixed in position above the skin surface and focused on the treatment site (Fig. 9). Thermal camera images were taken at 20 ms intervals starting at approximately 2.5 ms following a laser pulse. The image exposure time was approximately 2 ms. The imager sensor detects thermal light emitted from the surface and also some from the subsurface as light of wavelengths 3-5 μm has a penetration depth of about 50 μm in the tissue. The measured temperatures therefore represented a weighted average of the skin temperature within the penetration depth of the detected thermal radiation.

**Nd:YAG** 

**Laser** Skin

Fig. 9: Experimental set-up.

25

New Skin Treatment Possibilities with PIANO Mode on an Nd:YAG Laser

Skin surface measurements of the temperature 10

10

distribution following Nd:YAG and diode laser irradiations were performed in-vivo on female

8

8

patients’ of skin type II.

6

6

**)**

**)**

**C**

**C**

**0**

**0**

**(**

**(**

4

4

**T**

**T**

**III. RESULTS**

Δ

Δ

2

2

Figure 10 shows the measured temperature evolution on the hand skin (thenar web between

0

0

thumb and forefinger) following Nd:YAG laser radiation with a 1.5 s PIANO pulse (Super Long Pulse), and a 20 ms LP pulse (Long Pulse), for the

LP 20 ms

LP 20 ms

PIANO 1.5 s

PIANO 1.5 s

0 50 100 150 200

0 50 100 150 200 **Fluence (J/cm2)**

**Fluence (J/cm2)**

same pulse fluence of 35 J/cm2.

44

44

42

42

40

Fig. 11: Measured dependence of the peak skin surface temperature increase for the PIANO mode Nd:YAG laser illumination of 1.5 s pulse duration. For comparison, a temperature increase with a shorter, 20 ms Nd:YAG pulse is also shown. Spot sizes were in the range of 10-20 mm. Diagonal lines in the chart are only a visual aid.

) )

C

C

0

0

((

T

T

Δ

Δ

40 

38

38

36

36

34

34

0 10 20 30 40 50 60

0 10 20 30 40 50 60 Time (s)

Time (s)

As expected, the skin temperature increase depended on the body location. Table 1 shows the measured temperature increase on the facial cheek, beard and forehead. For comparison, the temperature increase on the hand skin is also shown.

Fig. 10: Measured temperature evolution of dorsal hand skin following irradiation with the Nd:YAG laser PIANO (1.5 s) pulse mode (red line), and LP (20 ms) pulse mode. The laser pulse fluence of 35 J/cm2 was the same for both pulse duration modes.

As expected from the calculations (See Fig. 5), the measured peak epidermal temperature is approximately 3 x higher with the 20 ms Nd:YAG laser pulse. After approximately 0.5 s when the epidermal temperature gets equalized with the temperature of the lower lying dermis the surface skin temperature has approximately the same temperature decay dependence, irregardless of the Nd:YAG pulse duration. This demonstrates that with the PIANO mode, the same temperatures of the bulk dermis as with shorter pulse durations, are achieved. However, with the PIANO mode, high initial temperature peaks in the epidermis are completely avoided.

Figure 11 shows the measured dependence of the measured peak skin surface (hand skin between thumb and forefinger) temperature on the illuminating PIANO pulse fluence.

The hand skin (thenar web) temperature increase was observed to follow an approximately linear dependence on the PIANO laser pulse fluence, independent of the pulse duration, in the range of 0.5 to 5 s.

**Table 1: Skin temperature increase following a PIANO pulse (1.5 s) illumination on different parts of the body. Measurements were performed on a female patient of skin type II.**

|  |  |  |
| --- | --- | --- |
|  | Δ**T** Δ**T** **(at 35 J/cm2) (at 35 J/cm2)** | Δ**T/Fluence** Δ**T/Fluence**  |
| cheek cheek  | 7.6 0 C 7.6 0 C | 0.22 0C cm2 /J 0.22 0C cm2 /J  |
| beard beard  | 7.1 0 C 7.1 0 C | 0.20 0C cm2 /J 0.20 0C cm2 /J  |
| forehead forehead  | 12.7 0 C 12.7 0 C | 0.36 0C cm2 /J 0.36 0C cm2 /J  |
| hand hand  | 2.7 0 C 2.7 0 C | 0.05 0C cm2 /J 0.05 0C cm2 /J  |

Typically, PIANO laser treatments are painless since the epidermis does not reach as high temperatures as is the case with shorter pulses. However, this is not the case when a laser treatment is performed on the skin area where there is an underlying larger size blood vessel. In such a case, the temperature of the blood vessel and of the neighboring tissue can substantially increase, resulting in the feeling of pain.

It is interesting to note that the presence of an underlying blood vessel can be detected from the temporal evolution of the skin surface temperature. As an example, Fig. 12 shows the temporal evolution of the skin surface temperature on dorsal hand skin where an underlying blood vessel was present.

26

New Skin Treatment Possibilities with PIANO Mode on an Nd:YAG Laser

) )

CC

0

0

((

T

T

Δ

Δ

Blood vessell 

Blood vessell

40

40

38

38

36

36

34

34

70

70

0 10 20 30 40 50 60

0 10 20 30 40 50 60 Time (s)

Time (s)

proteins (HSP). A controlled temperature increase and duration of the thermal surge within the dermal tissue is crucial to obtaining a suitable photo-thermal biologic response. Heating the collagen to a higher temperature than required to achieve a controlled heat shock would denature and damage the collagen beyond the point of contraction. There is a logarithmic relationship between duration and temperature increase that is required to achieve HSP expression. For heat shocks that last several seconds,

Fig. 12: Temporal evolution of dorsal hand skin in the presence of an underlying blood vessel following an Nd:YAG PIANO pulse (1.5 s, 150 J/cm2). Heat diffusion from the underlying heated vessel results in a second skin surface temperature peak, with the time delay depending on the depth of the vessel within the skin.

The underlying, more absorbant blood vessel gets heated up during the Nd:YAG laser irradiation. Following the laser pulse, the heat diffuses away from the hot blood vessel in all directions. With a certain time delay, this results in another temperature peak within the epidermis.

**IV. DISCUSSION**

**a) PIANO skin remodeling**

Beginning in our 20s, the effects of aging start to become visible in skin. Genetically programmed chronologic aging causes biochemical changes in collagen and elastin, the connective tissues that give skin its firmness and elasticity. As skin becomes less elastic, it also becomes drier. Underlying fat padding begins to disappear. With the loss of underlying support by fat padding and connective tissues, the skin begins to sag. It looks less supple, and wrinkles form. Occurring simultaneously with genetically programmed aging, the process of photoaging is also taking place. Photoaging is the effect of chronic and excessive sun exposure on the skin. Cigarette smoking also contributes to aging effects.

Laser remodeling is a procedure inducing controlled thermal injury to the collagen in the dermis which leads to subsequent collagen remodeling while preserving the epidermis [13]. As dermal collagen heals and remodels under an undamaged outer layer, the skin tightens and smoothes, thus improving its appearance.

The aim of the controlled, sub-coagulative laser heating of the dermis is to induce a heat shock response (HSR). HSR is defined as temporary changes in cellular metabolism. These changes are rapid and transient, and are characterized by the production of a small family of proteins termed the heat shock

HSP hyper-expression occurs at relatively low sub coagulative skin temperatures, between 50 and 55 0C. For shorter thermal surges the temperature elevations must be considerably higher [13].

Photo-thermal remodeling has been successfully achieved, reaching a critical sub-coagulative temperature, with many different lasers, radiofrequency devices and polychromatic high

intensity light sources [16-47]. Since for many of the successfully used devices, the pulse duration was on the order of seconds, this demonstrates that deep, homogeneous heating of the dermis plays an important role in photo-thermally activated neo

collagenesis. It has been demonstrated, however, that with extremely short and high power laser sources (Q switched and possibly FRAC3 Nd:YAG), an additional, photo-mechanical inflammatory effect comes into play [48]. The photo-thermal, homogeneous photo thermolysis promotes the formation of collagen type I, while the photo mechanical, selective photo-thermolysis promotes the synthesis of collagen type III more effectively.

There are several advantages of Nd:YAG skin remodeling compared to using other devices. First, the absorption in the epidermis is smaller compared to other light-based devices. Second, the depth of the thermal surge within the skin is completely controlled, and there is no danger of accidentally overheating deeper lying tissue, as is the case with radiofrequency devices. Third, the depth of the thermal surge can be adjusted to a certain extent by varying the laser spotsize. And fourth, the energy is delivered into the skin with energy-controlled pulses of adjustable duration.

Nd:YAG laser-induced collagen remodeling has been studied extensively. Goldberg first reported the effect of non-ablative laser skin resurfacing with the Q-switched 1,064-nm neodymium–yttrium–aluminum garnet (Nd:YAG) laser in 1997 [16]. They found that the Q-switched 1,064-nm Nd:YAG laser provides satisfactory clinical results with no post-operative morbidity in the treatment of periocular and perioral rhytides. Later Goldberg performed a series of clinical studies with histologic examinations to prove the collagen remodeling effects of a Q-switched 1,064-nm

27

New Skin Treatment Possibilities with PIANO Mode on an Nd:YAG Laser

Nd:YAG laser [17,18]. A long-pulsed 1,064-nm Nd:YAG laser with its low scattering coefficient and weak absorption by water and melanin has also been shown to improve the appearance of coarse wrinkles and fine lines and to reduce skin laxity [32–36]. In a recent systematic study by Liu et al on a Kunming (KM) mouse model in vivo, the following results were obtained with a long-pulsed (0.3 ms) Nd:YAG laser at 30 J/cm2 and 6 mm spotsize [48]. Eight weeks after the Nd:YAG laser treatments, an increase in collagen type I (46.4+-4.6%) and collagen type III (15.9+- 2.9%) were observed. Fibroblast proliferation correlated with collagen synthesis. The collagen was observed to begin to increase 3 weeks after the laser treatments, and the increasing trend could persist for 5 weeks, reaching its greatest level of effect in 4 weeks, a result similar to human clinical trials [48].

It is important to note that when only photo thermal, bulk heating of the dermis is indicated, the pulsewidth at which the Nd:YAG laser energy is delivered to the skin does not have a significant effect on the overall temporal evolution of the dermal temperature. No matter whether the laser energy is delivered in 0.3 ms or 500 ms, the dermal temperature will quickly, within a hundred milliseconds, equalize with the temperature of the epidermis, and then slowly cool down with its long relaxation time of approximately ten seconds. The laser-induced heat shock of the bulk dermis therefore always lasts for several seconds, irregardless of the pulsewidth of the heating Nd:YAG laser source. Shorter pulsewidths have a significant effect only on the bulk epidermis, which in the initial heating phase may experience a higher, short duration thermal surge due to its higher absorption compared to the dermis. For “heat shocking” of the dermis using Nd:YAG laser radiation, it is therefore preferable to use super long PIANO pulses that achieve approximately the same overall heat shock effect on the dermis while sparing the epidermis from the unnecessary thermal damage.

Based on the results of the previous and present study, we recommend the following starting treatment parameters for skin remodeling (Table 2).

**Table 2: Starting treatment parameters for PIANO mode Nd:YAG skin remodeling. Parameters are for the Fotona Dynamis Nd:YAG laser system.**

|  |
| --- |
| **SKIN REMODELING**  |
| SOURCE: Nd:YAG  | HANDPIECE: R33, R34  |
| MODE: PIANO |  |
| Pulse duration: 0.5 - 1.5 s  |  |
| Fluence: 35 – 65 J/cm2 |  |
| Spotsize: 9 – 20 mm |  |

**b) PIANO skin tightening**

The 1064 nm laser has been shown to be an excellent wavelength to contract collagen and initiate new collagen formation, resulting in what is commonly referred to as “skin tightening” [70-72]. The advantages of using the Nd:YAG to treat for skin laxity include its speed of treatment (due to large treatment spot sizes and high repetition rate), fluence levels (for sufficient heat deposition), low melanin absorption (to avoid epidermal injury) and cost (treatments can be delivered quickly, safely and without disposables). Since skin tightening belongs to the skin remodeling category it is a result of sufficient heat deposition to contract and reformulate collagen.

The before and after pictures after the Nd:YAG skin tightening treatments show a remarkable improvement in the patient’s skin color, tone, wrinkles and laxity, with minimal patient discomfort. Complications from the laser treatment were minimal including erythema and swelling from a few hours, to a few days post treatment. No other complications were seen. The treatment included minimal discomfort and high patient satisfaction [70].

Since a PIANO mode has been introduced only recently, the longest, until recently available pulse durations (50 ms) were used for skin tightening. Using a so called Taylor technique, [70] a relatively low single pulse fluence of 18 J/cm2 was applied in order to avoid unnecessary epidermal damage. The required dose of heat deposition was then achieved by repeatedly treating areas 3-4 times. The clinical endpoint of the laser treatment was a surface skin temperature of 40-45ºC as measured by IR thermometry.

As observed also by Taylor, [72] it is less important how quickly the necessary heat dose is introduced into the skin than ensuring sufficient energy has been delivered. The PIANO mode is thus perfectly suited for performing skin tightening treatments. Instead of treating a particular skin area 3-4 times with 50 ms pulses of low (18 J/cmm2) pulse fluence, the PIANO mode enables the doctor to tread the skin within one pass only, using a higher fluence of 35-65 J/cm2 (See Table 2).

**c) PIANO wound healing and scar prevention**

The cosmetic outcome of surgical scars is of paramount importance to physicians and patients undergoing surgery. Considerable efforts are therefore made to improve scar appearance and, more importantly, to avoid the development of postsurgical hypertrophic scars or keloids. Prevention of hypertrophic scars is obviously preferable to treatment

28

New Skin Treatment Possibilities with PIANO Mode on an Nd:YAG Laser

and implies using a therapy aimed at reducing their incidence [49, 50].

Recently, a significant improvement in the appearance of a surgical scar was obtained by performing a LASH (Laser-Assisted Skin Healing) treatment immediately after surgery. Initially, Capon et al [61] showed the ability of an 810-nm diode-laser system to assist in wound closure. Acceleration of wound healing and an indiscernible scar were obtained in hairless rats. This finding was later confirmed in a pilot study of five patients. Laser-treated scar portions demonstrated better quality compared with untreated scar portions [62]. A study was also reported that showed for the first time the possibility of improving the appearance of hypertrophic scars by altering, through a controlled thermal stress, the wound-healing process immediately after conventional hypertrophic scar revision [63]. Finally, a prospective comparative clinical study was made to evaluate the use of an 810 nm diode laser system to accelerate and improve the healing process in surgical scars immediately after skin closure [69]. Twenty-nine women and 1 man (mean age = 41.4 years; Fitzpatrick skin types I-IV) were included to evaluate the safety and performance of the laser system. The laser dose (or fluence in J/cm2) was selected as a function of phototype and skin thickness. Each surgical incision (e.g., abdominoplasty) was divided into two parts. An 8-cm segment was treated

The above LASH (Laser-Assisted Skin Healing) technique induces a temperature elevation in the skin which modifies the wound-healing process. As in laser remodeling, skin temperature elevation plays a major role in modifying the wound-healing process. As previously demonstrated in experimental evaluations on animals, temperature elevation results in a marked increase in levels of heat shock protein 70 (HSP70) in skin structures, particularly around blood vessels, hair follicles, and sebaceous glands [64]. Heat shock response (in particular HSP70 synthesis) is responsible for the release and production of growth factors (in particular via the modification of the TGFb profile), thus increasing the rate of cell proliferation and the speed of collagen production, hence improving scar aspects [65, 66]. Several studies have demonstrated that the most significant difference between normal tissue and scar tissue is in the orientation of the fibrosis matrix [67, 68].

In order to compare the thermal effects of the 810 nm diode laser, used in the LASH technique, with the PIANO mode Nd:YAG laser, we have made measurements of hand skin (thenar web between thumb and forefinger) surface temperature following illumination with a diode laser (810nm). Results are shown in Fig. 13.

10

10

10

10

10

Diode 810 nm ( 5 s)

Diode 810 nm ( 5 s)

Diode 810 nm ( 5 s)

Diode 810 nm ( 5 s)

9

9

9

9

9

with the laser immediately after skin closure. A 8

8

8

8

8

separate 8-cm segment was left untreated as a control. 7

7

7

7

7

Clinical evaluations (overall appearance ratings, 6

6

6

6

6

**)**

**)**

**)**

**)**

**)**

comparative scar scale) of all scars were conducted at

5

5

5

5

5

**C**

**C**

**C**

**C**

**C**

**0**

**0**

**0**

**0**

**0**

**(**

**(**

**(**

**(**

**(**

4

4

4

4

4

10 days, 3 months, and 12 months by both surgeon

**T**

**T**

**T**

**T**

**T**

Δ

Δ

Δ

Δ

Δ

3

3

3

3

3

and patients. Profilometry analysis from silicone 2

2

2

2

2

replicas of the skin was done at 12 months. Wilcoxon 1

1

1

1

1

signed-rank test analyses were performed. Twenty-two 0

0

0

0

0

patients were treated using a high dose (80–130 J/cm2) and 8 patients with a low dose (<80 J/cm2). At 12 months in the high-dose group, both surgeon

Diode 810 nm ( 5 s)

0 50 100 150 200 0 50 100 150 200 0 50 100 150 200 0 50 100 150 200 0 50 100 150 200 **Fluence (J/cm2)**

**Fluence (J/cm2)**

**Fluence (J/cm2)**

**Fluence (J/cm2)**

**Fluence (J/cm2)**

and patients reported an improvement rate of the laser treated segment over the control area of 72.73 and 59.10%, respectively. For these patients, profilometry results showed a decrease in scar height of 38.1% (p = 0.027) at 12 months for the laser-treated segment versus the control. Three patients treated with higher doses (115 J/cm2) experienced superficial burns on the laser-treated segment, which resolved in about 5–7 days. For the eight patients treated at low dosage (<80 J/cm2), there was no significant difference in the treated segment versus the control segment. No side effects were observed. In conclusion, this prospective comparative trial demonstrated that an 810-nm diode laser treatment, performed immediately after surgery, can improve the appearance of a surgical scar. The dose plays a significant role in scar improvement and must be well controlled.

Fig. 13: Measured dependence of the peak skin surface temperature increase for diode (810 nm) laser illumination with a duration of 5 s. Spot size was 6 mm. The diagonal line in the chart is only a visual aid.

When comparing the results obtained with the diode laser (Fig. 13) and the PIANO mode Nd:YAG laser (See Fig. 11) we can conclude that the PIANO mode Nd:YAG laser is suitable for the LASH technique, with similar fluences required to induce the recommended temperature elevation in the skin, and thus to modify the wound-healing process.

Based on the results of the previous and present study, we recommend the following starting treatment parameters for Nd:YAG laser-wound healing and scar prevention (Table 3).

29

New Skin Treatment Possibilities with PIANO Mode on an Nd:YAG Laser

**Table 3: Starting treatment parameters for PIANO mode Nd:YAG promotion of wound healing and scar prevention. Parameters are for the Fotona Dynamis Nd:YAG laser system.**

|  |
| --- |
| **WOUND HEALING/SCAR PREVENTION**  |
| SOURCE: Nd:YAG  | HANDPIECE: R33, R34  |
| MODE: PIANO |  |
| Pulse duration: 5 s  |  |
| Fluence: 90 – 110 J/cm2 |  |
| Spotsize: 3 – 20 mm |  |

**V. CONCLUSIONS**

In conclusion, a new Nd:YAG laser pulse duration modality was introduced. This new, super long, PIANO modality extends the Nd:YAG pulse durations to the seconds regime. The duration of the PIANO mode is much longer than the thermal relaxation time of the epidermis, or any other skin structures, such as hair follicles or blood vessels. As a result, the PIANO modality does not cause high initial temperature peaks in the epidermis. For “heat shocking” of the dermis using Nd:YAG laser radiation, it is therefore preferable to use super long PIANO pulses that achieve approximately the same overall heat shock effect on the dermis while sparing the epidermis from the unnecessary thermal damage. The PIANO modality is thus perfectly indicated for treatments where overall homogeneous, bulk heating of the dermis is desired.

Clinical protocol guidelines for skin remodeling, skin tightening and wound healing were developed, based on previously published clinical studies and our in-vivo study of the thermal effects of the PIANO mode on human skin.

The PIANO modality may be of benefit also when considering treatments that combine the non-ablative Nd:YAG laser skin remodeling with the Er:YAG laser fractional skin resurfacing [73-74].

**REFERENCES**

1. Goldman L, Rockwell RJ Jr. Laser Systems and their Application in Medicine and Biology. Adv Biomed Eng Med Phys., 1968; 1:317-82.

2. Anderson RR, Parrish JA. Selective Photothermolysis-precise Microsurgery by Selective Absorption of Pulsed Radiation. Science, 1983; 220(4596):524-7.

3. Svaasand IO, Milner TE, Anvari B, Norvang LT, Tanenbaum BS, Kimel S, Berns MW, Nelson JS. Epidermal Heating During Laser Induced Photothermolysis of Port Wine Stains: Modeling Melanosomal Heating after Dynamic Cooling the Skin Surface, Laser Interaction with Hard and Soft Tissue II. SPIE Proc., 1995; 2323:366-77.

4. Nakagawa H, Tan OT, Parish JA. Ultrastructure Changes in Human Skin after Exposure to a Pulsed Laser. J Invest Dermatol., 1985; 84:396-400.

5. Anderson R, Parrish J. The optics of human skin. J Invest Dermatol., 1981; 77:13-9.

6. Lukac M, Grad L. Scanner Optimized Aesthetic Treatments with the VSP Nd:YAG Lasers. J Cosmet Laser Ther, June 008; Vol. 10, No. 2.

7. Lukac M, Zabkar J, Gorjan M, Vizintin Z. Beyond Customary Paradigm: Self-Induced Fractional Nd:YAG Laser Hair Removal. J Laser Health Academy, 2010; 1:35-46,

www.laserandhealth.com.

8. Lukac M, Sult T, Zabkar J, Gorjan M, Vizintin Z. Parameters for the New FRAC3 Nd:YAG laser skin treatment modality. J Laser Health Academy, 2010; 1:47-55,

www.laserandhealth.com.

9. Cencic B, Lukac M, Zabkar J, Marincek M, Vizintin Z. High fluence, high beam quality Q-switched Nd:YAG laser with optoflex delivery system for treating benign pigmented lesions and tattoos. J Laser Health Academy, 2010; 1:9-18, www.laserandhealth.com.

10. PIANO Nd:YAG laser mode was developed by Fotona d. d., www.fotona.com.

11. www.meriam-webster.com.

12. Capon A, Mordon S. Can Thermal Lasers Promote Skin Wound Healing? Am J Clin Dermatol., 2003; 4(1):1-12. 13. Goldberg DJ, Whitworth J. Laser skin resurfacing with the Q-switched Nd:YAG laser. Dermatol Surg., 1997, 23(10): 903-6 .

14. Goldberg DJ. Q-switched Nd:YAG laser: Rhytid improvement by non-ablative dermal remodeling, J Cutan Laser Ther., 2000; 2(3):157-60.

15. www.fotona.com.

16. Goldberg DJ, Silapunt S. Histologic evaluation of a Q-switched Nd:YAG laser in the non-ablative treatment of wrinkles. Dermatol Surg., 2001; 27(8):744–46.

17. Lask GL, Lee PK, Seyfadeh M, et al. Non-ablative laser treatment of facial rhytids. SPIE Proc, 1997; 2970:338–49. 18. Trelles MA, Allones I, Luna R. Facial rejuvenation with a non-ablative 1320 nm Nd:YAG laser: A preliminary clinical and histologic evaluation. Dermatol Surg., 2001; 27:111–6. 19. Levy JL, Trelles M, Lagarde JM, Borrel MT, Mordon S. Treatment of wrinkles with the non-ablative 1,320-nm Nd:YAG laser. Ann Plast Surg, 2001 Nov; 47(5):482-8. 20. Alster TS, Kurban AK, Grove GL, Tan OT. Alteration of argon laser-induced scars by the pulsed dye laser. Lasers Surg Med., 1993; 13(3):368-73.

21. Alster TS. Improvement of erythematous and hypertrophic scars by the 585 nm flashlamp-pumped pulsed dye laser. Ann Plastic Surg., 1994 Feb; 32(2):186-90.

22. Alster TS, William CM. Treatment of keloid sternotomy scars with 585 nm flashlamp-pumped pulsed dye laser. Lancet, 1995 May 13; 345(8959):1198-200.

23. Zelickson BD, Kilmer SL, Bernstein E, Chotzen VA, Dock J, Mehregan D, Coles C. Pulsed dye laser therapy for sun damaged skin. Lasers Surg Med., 1995; 25(3):229–36. 24. Zelickson B, Kist D. Effect of pulsed dye laser and intense pulsed light source on the dermal extracellular matrix remodeling. Lasers Surg Med., 2000; 12(Suppl):17. 25. Tanghetti EA, Sherr EA, Alvarado SL. Multipass treatment of photodamage using the pulse dye laser. Dermatol Surg., 2003; 29:686–91.

26. Dahiya R, Lam S, Williams EF. A systematic histologic analysis of non-ablative laser therapy in a porcine model using the pulse-dye laser. Arch Facial Plast Surg., 2003; 5(3):218–23. 27. Goldberg DJ, Sarradet D, Hussain M, Krishtul A, Phelps R. Clinical, histologic and ultrastructural changes after non ablativetreatment with a 595-nm flashlamp-pumped pulsed dye laser: Comparison of varying settings. Dermatol Surg.,

30

 2004; 30:979–82.

New Skin Treatment Possibilities with PIANO Mode on an Nd:YAG Laser follow-up study. Arch Dermatol., 2004; 140(11):1337–41.

28. Hsu TS, Zelickson B, Dover JS, Kilmer S, Burns J, Hruza G. Multicenter study of the safety and efficacy of a 585 nm pulsed-dye laser for the non-ablative treatment of facial rhytides. Dermatol Surg., 2005; 31(1):1–9.

29. Rogachefsky AS, Becker K, Weiss G, Goldberg DJ. Evaluation of a long-pulsed Nd:YAG laser at different parameters: An analysis of both fluence and pulse duration. Dermatol Surg., 2002; 28:932–35.

30. Dayan S, Damrose JF, Bhattacharyya TK, Mobley SR, Patel MK, O’Grady K, Mandrea S. Histological evaluations following 1,064-nm Nd:YAG laser resurfacing. Lasers Surg Med., 2003; 33(2):126–31.

31. Dayan S, Vartanian AJ, Menaker G. Non-ablative laser resurfacing using a long pulse 1064-nm Nd:YAG laser. Arch Facial Plast Surg., 2003 Jul-Aug; 5:310–5.

32. Schmults CD, Phelps R, Goldberg DJ. Non-ablative facial remodeling: Erythema reduction and histologic evidence of new collagen formation using a 300-microsecond 1064-nm

 Nd:YAG laser. Arch Dermatol., 2004 Nov; 140(11):1373–6. 33. Trelles MA, Alvarez X, Martin-Vazquez MJ, Trelles O, Velez M, Levy JL, Allones I. Assessment of the efficacy of Non-ablative long-pulsed 1064-nm Nd:YAG laser treatment of wrinkles compared at 2, 4, and 6 months. Facial Plast Surg., 2005; 21(2):145–53.

34. Anderson RR, Parrish JA. Selective photothermolysis: Precise microsurgery by selective absorption of pulsed radiation. Science, 1983; 220:524–7.

35. Fatemi A, Weiss MA, Weiss RA. Short-term histologic effects of non-ablative resurfacing: Results with a dynamically cooled millisecond-domain 1320 nm Nd:YAG laser. Dermatol Surg., 2002; 28(2):172–6.

36. Sadick N, Schecter AK. Utilization of the 1320-nm Nd:YAG laser for the reduction of photoaging of the hands. Dermatol Surg., 2004; 30(8):1140-4.

37. Fluck M, Giraud M-N, Tunc V, Chiquet M. Tensile stress dependent collagen XII and fibronectin production by fibroblasts requires separate pathways. Biochim Biophys Acta., 2003 Feb 17; 1593(2-3):239–48.

38. Varani J, Schuger L, Dame MK, Leonard C, Fligiel SE, Kang S, Fisher GJ, Voorhees JJ. Reduced fibroblast interaction with intact collagen as a mechanism for depressed collagen

synthesis in photodamaged skin. J Invest Dermatol., 2004; 122(6):1471–9.

39. Cook H, Davies KJ, Harding KG, Thomas DW. Defective extracellular matrix reorganization by chronic wound fibroblasts is associated with alterations in TIMP-1, TIMP-2, and MMP-2 activity. J Invest Dermatol., 2000;115(2):225-33.

40. Moseley R, Stewart JE, Stephens P, Waddington RJ, Thomas DW. Extracellular matrix metabolites as potential biomarkers of disease activity in wound fluid: Lessons learned from other inflammatory diseases? Br J Dermatol., 2004; 150(3):401–13. 41. Baumann L, Kaufman J, Saghari S. Collagen fillers. Dermatol Ther., 2006; 19(3):134–40.

42. Varani J, Dame MK, Rittie L, Fligiel SE, Kang S, Fisher GJ, Voorhees JJ. Decreased collagen production in chronologically aged skin: Roles of age-dependent alteration in fibroblast function and defective mechanical stimulation.

 Am J Pathol., 2006; 168(6):1861–8.

43. Lovell CR, Smolenski KA, Duance VC, Light ND, Young S, Dyson M. Type I and III collagen content and fibre distribution in normal human skin during ageing. Br J Dermatol., 1987; 117(4):419–28.

44. Fernandes D. Minimally invasive percutaneous collagen induction. Oral Maxillofac Surg Clin North Am., 2005; 17:51–63 vi.

45. Friedman PM, Jih MH, Skover GR, Payonk GS, Kimyai- Asadi A, Geronemus RG. Treatment of atrophic facial acne scars with the 1064-nm Q-switched Nd:YAG laser: Six-month

46. Yaghmai D, Garden JM, Bakus AD, Massa MC. Comparison of a 1,064 nm laser and a 1,320 nm laser for the nonablative treatment of acne scars. Dermatol Surg., 2005; 31(8 Pt 1):903–9.

47. Alam M, Hsu TS, Dover JS, Wrone DA, Arndt KA. Non ablative laser and light treatments: Histology and tissue effects-a review. Lasers Surg Med., 2003; 33:30–9. 48. Liu H, Dang Y, Wang Z, Chai X, Ren Q, Laser Induced Collagen Remodeling: a Comparative Study in vivo on Mouse Model. Lasers Surg Med., 2008; 40:13-9.

49. Rhett JM, Ghatnekar GS, Palatinus JA, O’Quinn M, Yost MJ, Gourdie RG. Novel therapies for scar reduction and regenerative healing of skin wounds. Trends Biotechnol, 2008; 26:173–80.

50. Mustoe TA, Cooter RD, Gold MH, Hobbs FD, Ramelet AA et al. International clinical recommendations on scar management. Plast Reconstr Surg., 2002; 110:560–71.

51. Tierney E, Mahmoud BH, Srivastava D, Ozog D, Kouba DJ Treatment of surgical scars with non-ablative fractional laser versus pulsed dye laser: a randomized controlled trial. Dermatol Surg., 2009; 35(8):1172–80.

52. Colwell AS, Longaker MT, Lorenz HP. Fetal wound healing. Front Biosci., 2003; 8:1240–8.

53. West TB. Laser resurfacing of atrophic scars. Dermatol Clin., 1997; 15:449–57.

54. Alster TS, West TB. Treatment of scars: a review. Ann Plast Surg., 1997; 39:418–32.

55. Chan HH, Wong DS, Ho WS, Lam LK, Wei W. The use of pulsed dye laser for the prevention and treatment of hypertrophic scars in Chinese persons. Dermatol Surg., 2004; 30:987–94, discussion 994.

56. Kye YC. Laser therapy of skin diseases. Korean J Dermatol., 2003; 41(1):1–6.

57. Nouri K, Jimenez GP, Harrison-Balestra C, Elgart GW. 585-nm pulsed dye laser in the treatment of surgical scars starting on the suture removal day. Dermatol Surg., 2003; 29:65–73.

58. Conologue TD, Norwood C (2006) Treatment of surgical scars with the cryogen-cooled 595 nm pulsed dye laser starting on the day of suture removal. Dermatol Surg., 2006; 32:13–20.

59. Choe JH, Park YL, Kim BJ, Kim MN, Rho NK et al. Prevention of thyroidectomy scar using a new 1,550-nm fractional erbium-glass laser. Dermatol Surg., 2009; 35(8):1199–205.

60. Capon A, Souil E, Gauthier B, Sumian C, Bachelet M et al. Laser-assisted skin closure (LASC) by using a 815-nm diode laser system accelerates and improves wound healing. Lasers Surg Med., 2001; 28:168–75.

61. Capon AC, Gosse AR, Iarmarcovai GN, Cornil AH, Mordon SR. Scar prevention by laser-assisted scar healing (LASH): a pilot study using an 810-nm diode-laser system. Lasers Surg Med., 2008; 40:443–5.

62. Capon A, Iarmarcovai G, Mordon S. Laser-assisted skin healing (LASH) in hypertrophic scar revision. J Cosmet Laser Ther, 2009; 11(4):220–3.

63. Souil E, Capon A, Mordon S, Dinh-Xuan AT, Polla BS, Bachelet M. Treatment with 815-nm diode laser induces long lasting expression of 72-kDa heat shock protein in normal rat skin. Br J Dermatol., 2001; 144:260–6.

64. Capon A, Mordon S. Can thermal lasers promote skin wound healing? Am J Clin Dermatol., 2003; 4:1–12.

65. Shah M, Revis D, Herrick S, Baillie R, Thorgeirson S et al. Role of elevated plasma transforming growth factor-beta1 levels in wound healing. Am J Pathol., 1999; 154:1115–24.

66. Dallon J, Sherratt J, Maini P, Ferguson M. Biological implications of a discrete mathematical model for collagen deposition and alignment in dermal wound repair. Math

31

New Skin Treatment Possibilities with PIANO Mode on an Nd:YAG Laser

Med Biol (formerly IMA J of Mathematics applied in medicine and biology), 2000; 17:379–93.

67. Dallon JC, Sherratt JA, Maini PK) Modeling the effects of transforming growth factor-beta on extracellular matrix alignment in dermal wound repair. Wound Repair Regen., 2001; 9(4):278–86.

68. Alam M, Pon K, Van Laborde S, Kaminer MS, Arndt KA, Dover JS. Clinical effect of a single pulsed dye laser treatment of fresh surgical scars: randomized controlled trial. Dermatol Surg., 2009; 32(1):21–5.

69. Capon A, Iarmarcovai G, Gonnelli D, Degardin N, Magalon G, Mordon S. Scar Prevention Using Laser Assisted Skin Healing (LASH) in Plastic Surgery. Aesthetic Plast Surg., 2010; 34(4): 438-46.

70. Taylor M, Gentle YAG Skin Tightening, Update 2009, Candela Clinical Bulletin No. 21.

71. Key D. Single-Treatment Skin Tightening By Radiofrequency and Long-Pulsed, 1064 nm Nd:YAG Laser Compared. Lasers Surg Med., 2007;39(2):169-75.

72. Taylor MB, Prokopenko I. Split-face comparison of radiofrequency versus long-pulse Nd:YAG treatment of facial laxity. J Cosmet Laser Ther, 2006; 8(1):17-22.

73. Lukac M, Perhavec T, Nemes K, Ahcan U. Ablation and thermal depths in VSP Er:YAG laser skin resurfacing. J Laser Health Academy, 2010; 1:56-71.

74. Marini L. SPF-RR Sequential Photothermal Fractional Resurfacing and Remodeling with the Variable Pulse Er:YAG and Scanner-Assisted Nd:YAG Laser. J Cosmet Laser Ther, 2009; 11(4):202-11.

The intent of this Laser and Health Academy publication is to facilitate an exchange of information on the views, research results, and clinical experiences within the medical laser community. The contents of this publication are the sole responsibility of the authors and may not in any circumstances be regarded as official product information by the medical equipment manufacturers. When in doubt please check with the manufacturers whether a specific product or application has been approved or cleared to be marketed and sold in your country.

32